Select OPTION NAME: XPD MAIN Kernel Installation & Distribution System

Select Kernel Installation & Distribution System <TEST ACCOUNT> Option: INSTallation

Select Installation <TEST ACCOUNT> Option: LOAD a Distribution

Enter a Host File: SRVR:[LAB]LAB\_AUTORELEASE\_1\_0\_V7.KID

KIDS Distribution saved on Mar 15, 2016@15:52:12

Comment: LAB AUTO-RELEASE 1.0 V7

This Distribution contains Transport Globals for the following Package(s):

Build LA\*5.2\*88 has been loaded before, here is when:

LA\*5.2\*88 Install Completed

was loaded on Jun 03, 2015@15:12:27

LA\*5.2\*88 Install Completed

was loaded on Jun 16, 2015@13:13:03

LA\*5.2\*88 Install Completed

was loaded on Jul 08, 2015@19:23:31

LA\*5.2\*88 Install Completed

was loaded on Sep 17, 2015@13:15:29

LA\*5.2\*88 Install Completed

was loaded on Nov 20, 2015@14:07:53

LA\*5.2\*88 Install Completed

was loaded on Feb 01, 2016@13:45:35

OK to continue with Load? NO// YES

Build LR\*5.2\*458 has been loaded before, here is when:

LR\*5.2\*458 Install Completed

was loaded on Jun 03, 2015@15:12:26

LR\*5.2\*458 Install Completed

was loaded on Jun 16, 2015@13:13:02

LR\*5.2\*458 Install Completed

was loaded on Jul 08, 2015@19:23:30

LR\*5.2\*458 Install Completed

was loaded on Sep 17, 2015@13:15:28

LR\*5.2\*458 Install Completed

was loaded on Nov 20, 2015@14:07:53

LR\*5.2\*458 Install Completed

was loaded on Feb 01, 2016@13:45:36

OK to continue with Load? NO// YES

Distribution OK!

Want to Continue with Load? YES//

Loading Distribution...

Build LA\*5.2\*88 has an Environmental Check Routine

Want to RUN the Environment Check Routine? YES//

LA\*5.2\*88

Will first run the Environment Check Routine, LA88A

Sending transport global loaded alert to mail group G.LMI

--- Environment is okay ---

LR\*5.2\*458

Use INSTALL NAME: LA\*5.2\*88 to install this Distribution.

Transport global for patch LA\*5.2\*88 loaded on Mar 16, 2016@11:56:59

Enter "VA to jump to VIEW ALERTS option

Select Installation <TEST ACCOUNT> Option: VERify Checksums in Transport Global

Select INSTALL NAME: LA\*5.2\*88 Loaded from Distribution 3/16/16@11:56:53

=> LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

This Distribution was loaded on Mar 16, 2016@11:56:53 with header of

LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

It consisted of the following Install(s):

LA\*5.2\*88 LR\*5.2\*458

Want each Routine Listed with Checksums: Yes// YES

DEVICE: HOME// ;80;66 VIRTUAL TELNET

PACKAGE: LA\*5.2\*88 Mar 16, 2016 11:57 am PAGE 1

-------------------------------------------------------------------------------

LA7UCFG Calculated 137316017

LA7UCFG1 Calculated 29526912

LA7UIO1 Calculated 75417661

LA7UTILB Calculated 23365722

LA7VHL Calculated 34933144

LA7VHLU8 Calculated 54853900

LA7VHLU9 Calculated 33364706

LA7VIN Calculated 31997123

LA7VIN1 Calculated 59376205

LA7VIN2 Calculated 46992282

LA7VIN2A Calculated 34119950

LA7VIN4 Calculated 81120393

LA7VIN4A Calculated 22773812

LA7VIN5 Calculated 84396228

LA7VIN5A Calculated 34198601

LA7VORC Calculated 22779822

LA88 Calculated 37005513

LA88A Calculated 100793409

18 Routines checked, 0 failed.

PACKAGE: LR\*5.2\*458 Mar 16, 2016 11:57 am PAGE 1

-------------------------------------------------------------------------------

LR458 Calculated 3643881

LRGP2 Calculated 21346835

LRLISTPS Calculated 17691343

LRNIGHT Calculated 7575682

LRVER5 Calculated 153025645

LRVR3 Calculated 91573916

LRVRAR Calculated 67819881

LRVRARU Calculated 32021539

8 Routines checked, 0 failed.

You have PENDING ALERTS

Enter "VA to jump to VIEW ALERTS option

Select Installation <TEST ACCOUNT> Option: BACKup a Transport Global

Select INSTALL NAME: LA\*5.2\*88 Loaded from Distribution 3/16/16@11:56:

53

=> LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

This Distribution was loaded on Mar 16, 2016@11:56:53 with header of

LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

It consisted of the following Install(s):

LA\*5.2\*88 LR\*5.2\*458

Subject: Backup of LA\*5.2\*88 install on Mar 16, 2016

Replace

Loading Routines for LA\*5.2\*88..................

Loading Routines for LR\*5.2\*458

Routine LR458 is not on the disk.........

Send mail to: LRUSER,DRI// LRUSER,DRI

Select basket to send to: IN//

And Send to:

Select Installation <TEST ACCOUNT> Option: INSTall Package(s)

Select INSTALL NAME: LA\*5.2\*88 Loaded from Distribution 3/16/16@11:56:

53

=> LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

This Distribution was loaded on Mar 16, 2016@11:56:53 with header of

LAB AUTO-RELEASE 1.0 V7 ;Created on Mar 15, 2016@15:52:12

It consisted of the following Install(s):

LA\*5.2\*88 LR\*5.2\*458

Checking Install for Package LA\*5.2\*88

Will first run the Environment Check Routine, LA88A

Sending install started alert to mail group G.LMI

Is this site using the Lab UI V1.6 interface? No// ?

Enter either 'Y' or 'N'.

Contact your Laboratory Information Manager to confirm.

**Is this site using the Lab UI V1.6 interface? No// YES**

Has the Lab UI COTS driver been upgraded to send HL7 v2.5.1 messages? No// ?

Enter either 'Y' or 'N'.

This normally involves a driver update on the COTS system

to allow the COTS system to send HL7 messages indicating

either HL7 v2.2 or v2.5.1. Contact your Laboratory Information

Manager to confirm the status of the driver update.

**Has the Lab UI COTS driver been upgraded to send HL7 v2.5.1 messages? No// YES**

Disabling Option [LA7 MAIN MENU]

Shutting down currently running Lab HL7 processes

Acquiring locks ...

Locks acquired.

N O T E: If you abort this installation

D RESTORE^LA88A from this console.

--- Environment is okay ---

Install Questions for LA\*5.2\*88

Incoming Files:

62.4 AUTO INSTRUMENT (Partial Definition)

Note: You already have the 'AUTO INSTRUMENT' File.

62.485 LA7 MESSAGE LOG BULLETINS (including data)

Note: You already have the 'LA7 MESSAGE LOG BULLETINS' File.

I will OVERWRITE your data with mine.

Checking Install for Package LR\*5.2\*458

Install Questions for LR\*5.2\*458

Incoming Files:

68.2 LOAD/WORK LIST (Partial Definition)

Note: You already have the 'LOAD/WORK LIST' File.

Want KIDS to Rebuild Menu Trees Upon Completion of Install? NO//

Want KIDS to INHIBIT LOGONs during the install? NO//

Want to DISABLE Scheduled Options, Menu Options, and Protocols? YES// NO

Enter the Device you want to print the Install messages.

Enter a '^' to abort the install.

DEVICE: HOME// VIRTUAL TELNET

LR\*5.2\*458



Enabling Option [LA7 MAIN MENU]

Be sure to restart the Lab Universal Interface Auto Download Job

\*\*\* Post install completed \*\*\*

Sending install completion alert to mail group G.LMI

Updating Routine file...

Updating KIDS files...

LR\*5.2\*458 Installed.

Mar 16, 2016@11:59:05

Not a production UCI

NO Install Message sent





100%  25 50 75 

Complete 

Install Completed