[**93430: SQATS - Release 0.5 - Verify image viewer ability to display**](https://DNS.3A93430)



**[DICOM multi-frame cine viewer](https://DNS)**

State:Draft

Originator:Gowru, Bhavana (LongView) Owner:Unassigned

Type:Manual

Test Data:Unassigned

Copied To:[SQATS - Build 3 - Verify image viewer ability to display DICOM multi-frame cine viewer](https://DNS.3A116536)

Description:

**Summary**

**Categories**

**Formal Review**

General Comments

**Manual Steps**

**Step 1**

**Execution Step**

Description

when user selects patient which have multiple DICOM images

Expected Results

Image viewer support DICOM multi-frame cine viewer

Comments

Validates

Attachments

**Step 2**

**Execution Step**

Description

Verify when user selects DICOM multi-frame image

Expected Results

Appropriate toolbar displays on the viewer

Comments

Validates

Attachments

**Step 3**

**Execution Step**

Description

All the images will be loaded into the viewer

Expected Results

All the frames will be played at a specified frame rate

Comments

Validates

Attachments

**Step 4**

**Execution Step**

Description

User should able to click on play button

Expected Results

Frames started to display with a specified frame rate

Comments

Validates

Attachments

**Step 5**

**Execution Step**

Description

User should able to stop the image display

Expected Results

Images will be paused

Comments

Validates

Attachments

**Step 6**

**Execution Step**

Description

When images are playing the text on image should display image number

Expected Results

check if the image number is changing when image is playing

Comments

Validates

Attachments

**Step 7**

**Execution Step**

Description

When the images are playing FPS value is displayed

Expected Results

Value should be displayed and changed according to specification

Comments

Validates

Attachments

**Step 8**

**Execution Step**

Description

User will be able to adjust the image viewing speed

Expected Results

TBD - how to adjust the frame rate

Comments

Validates

Attachments

**Step 9**

**Execution Step**

Description

Click on repeat series on the dicom toolbar

Expected Results

user should able to repeat the the series

Comments

Validates

Attachments

**Step 10**

**Execution Step**

Description

once user clicks on repeat series, toolbar option changed to repeat study

Expected Results

user should be able to repeat the study

Comments

I ~~VoUdol"~~

Attachments

**Associated E-Signatures**

Signed **Action Signer Comment Additional Information**